## Foglio1
|  | LINEARIZATION: |  |  | SBE335, serial number: 0059, 26/02/2013, INST vs. SPRT (Rosemount 162CE) + Hart Scientific 1593 SuperThermometer + Hart Scientific 7052 Seawater Bath
T_BATH (SPRT) | AVE RAW REF ZERO | AVE RAW REF RESISTOR FULL SCALE | AVE RAW THERMISTOR | (Max - Min) RAW REF ZERO | (Max - Min) RAW REF RESISTOR FULL SCALE | (Max - Min) RAW THERMISTOR | AVE RAW THERMISTOR CORRECTED FOR ZERO & FULL SCALE REF | CNV THERMISTOR | CNV THERMISTOR - T_BATH (SPRT)
ITS-90, °C | [7 empty] | ITS-90, °C | ITS-90, °C
27.9968 | 196.5293 | 1047524.7586 | 217264.2379 | 14.2414 | 18.3103 | 15.069 | 217326.3241 | 27.9972 | 0.00035455
27.9968 | 192.2467 | 1047502.7333 | 217257.0133 | 14.4 | 18.1 | 14.8 | 217327.0667
25.0359 | 196.3165 | 1047496.9677 | 243116.2903 | 14.9677 | 17.6452 | 14.9677 | 243215.7935 | 25.0358 | -9.5355e-05
25.0357 | 194.7638 | 1047499.2069 | 243115.1241 | 13.9655 | 17.7241 | 14.2759 | 243215.3069
20.1277 | 198.4531 | 1047499.7931 | 294111.9517 | 14.2069 | 18.0345 | 15.2069 | 294271.2345 | 20.1276 | -0.00013486
20.1274 | 196.689 | 1047492.8387 | 294110.4226 | 14.5161 | 17.3548 | 14.9333 | 294272.91
15.2226 | 198.942 | 1047504.0333 | 357585.8933 | 14.1333 | 18.3667 | 15.4 | 357820.6467 | 15.2227 | 6.8e-05
15.2229 | 193.7903 | 1047498.6897 | 357578.7759 | 14.5172 | 17.8621 | 15.6897 | 357818.7379
10.2491 | 192.4914 | 1047505.8571 | 438275.1571 | 14.5357 | 17.75 | 16.3571 | 438610.8464 | 10.2488 | -0.00030382
10.2487 | 192.2568 | 1047502.7742 | 438274.2774 | 14.6129 | 17.7742 | 16 | 438611.3226
5.1457 | 196.969 | 1047503.6333 | 543071.99 | 13.8333 | 18.4667 | 17.8667 | 543532.9433 | 5.1455 | -0.00020223
5.1459 | 192.228 | 1047498.5 | 543066.1967 | 14.1667 | 17.7667 | 18.1333 | 543532.0767
1.9582 | 197.7371 | 1047505.1786 | 663733.9036 | 14.5 | 18.4286 | 19.1429 | 664339.9214 | 1.9582 | -3.2782e-05
1.9584 | 194.8965 | 1047498.9677 | 663727.4839 | 13.6774 | 16.9355 | 19.2903 | 664338.4452
|  | LINEARIZATION: |  |  | SBE335, serial number: 0058, 26/02/2013, INST vs. SPRT (Rosemount 162CE) + Hart Scientific 1593 SuperThermometer + Hart Scientific 7052 Seawater Bath |  |  |  |  | New Slope: 1.000002 | New Offset: 0.000405
T_BATH (SPRT) | AVE RAW REF ZERO | AVE RAW REF RESISTOR FULL SCALE | AVE RAW THERMISTOR | (Max - Min) RAW REF ZERO | (Max - Min) RAW REF RESISTOR FULL SCALE | (Max - Min) RAW THERMISTOR | AVE RAW THERMISTOR CORRECTED FOR ZERO & FULL SCALE REF | CNV THERMISTOR | CNV THERMISTOR - T_BATH (SPRT) | Corrected CNV THERMISTOR - T_BATH (SPRT) | Corrected CNV THERMISTOR - T_BATH (SPRT)
ITS-90, °C | [7 empty] | ITS-90, °C | ITS-90, °C | ITS-90, °C | ITS-90, °C
27.9968 | 197.894 | 1047503.1667 | 228820.92 | 13.4333 | 17.4667 | 14.7667 | 228900.41 | 27.9966 | -0.000258 | 27.9971 | 0.00020299
27.9968 | 198.1783 | 1047503.4 | 228822.2233 | 13.2 | 18.0667 | 13.9333 | 228901.4433
25.0359 | 194.2316 | 1047502 | 256387.9839 | 14.129 | 17.7097 | 14.3226 | 256504 | 25.0352 | -0.00068703 | 25.0356 | -0.00023196
25.0357 | 199.0393 | 1047500.7241 | 256391.0207 | 13.7931 | 18.3793 | 14.7931 | 256503.7
20.1277 | 197.1867 | 1047501.8333 | 310853.7933 | 13.9667 | 17.8 | 15.0667 | 311033.7267 | 20.1271 | -0.0006778 | 20.1275 | -0.00023255
20.1274 | 198.2647 | 1047498.8 | 310855.37 | 13.7667 | 17.6667 | 14.6 | 311035.43
15.2226 | 192.958 | 1047506.4333 | 378789.33 | 14.2333 | 18.2 | 14.6333 | 379052.7733 | 15.2222 | -0.00042303 | 15.2226 | 1.2411e-05
15.2229 | 193.0955 | 1047499.7931 | 378785.3862 | 13.9655 | 18.1034 | 15.1379 | 379051.1138
10.2491 | 196.0003 | 1047499.7586 | 465344.2207 | 14.0345 | 18.5172 | 15.8966 | 465713.231 | 10.2482 | -0.00083617 | 10.2486 | -0.00041068
10.2487 | 192.2568 | 1047502.7742 | 438274.2774 | 14.6129 | 17.7742 | 16 | 438611.3226
5.1457 | 197.0193 | 1047505.9667 | 578001.89 | 13.8333 | 17.8667 | 17.1333 | 578503.9233 | 5.1449 | -0.00077583 | 5.1454 | -0.00036054
5.1459 | 194.6532 | 1047503.1613 | 577998.4903 | 13.6129 | 17.9355 | 17.5484 | 578503.0935
1.9582 | 196.7962 | 1047502.2414 | 622683.1759 | 14.1724 | 18.3793 | 18.7241 | 623241.5931 | 1.9575 | -0.00069833 | 1.9579 | -0.00028942
1.9584 | 190.9883 | 1047499.8333 | 622678.14 | 13.8 | 17.9333 | 18.6 | 623240.3433
FIXED POINTS: SBE35, serial number 0059, 27/02/2013
Fixed Point Cell | Nominal Cell TEMP | Corrected Cell TEMP | Corrected AVE RAW THERMISTOR | SBE35 TEMP | SBE35 TEMP - Corrected Cell TEMP
ITS-90, °C | ITS-90, °C | ITS-90, °C |  | ITS-90, °C | ITS-90, °C
TPW | 0.01 | 0.009808 | 678359.4733 | 0.0098093 | 1.2667e-06
MPGa | 29.7646 | 29.7647 | 203381.3 | 29.7647 | -5.3333e-05
FIXED POINTS: SBE35, serial number 0058, 27/02/2013
Fixed Point Cell | Nominal Cell TEMP | Corrected Cell TEMP | Corrected AVE RAW THERMISTOR | SBE35 TEMP | SBE35 TEMP - Corrected Cell TEMP
ITS-90, °C | ITS-90, °C | ITS-90, °C |  | ITS-90, °C | ITS-90, °C
TPW | 0.01 | 0.009808 | 723746.96 | 0.0094034 | -0.00040463 | 0.0098084
MPGa | 29.7646 | 29.7647 | 214047.0548 | 29.7641 | -0.00064416 | 29.7645
"AS RECEIVED" TEST: SBE35 with tip bushings, serial numbers 0058 & 0059 vs. SPRT (Rosemount 162CE + Hart Scientific 1590 SuperThermometer + Hart Scientic 7052 Seawater Bath
T_BATH (SPRT) | SBE35-0058 | SBE35-0058 - T_BATH | SBE35-0059 | SBE35-0059 - T_BATH | SBE35-0059 - SBE35-0058
ITS-90, °C | ITS-90, °C | ITS-90, °C | ITS-90, °C | ITS-90, °C | ITS-90, °C
18.1706 | 18.1702 | -0.00032962 | 18.1704 | -0.0001352 | 0.00019443